Theory encounters experiments of Mössbauer spectroscopy
Tutorial lecture
Hartmut Spiering
Institut für Anorganische Chemie und Analytische Chemie,
Johannes Gutenberg Universität Mainz, Mainz, Germany
Dénes Lajos Nagy
Wigner Research Centre for Physics, Hungarian Academy of Sciences,
Budapest, Hungary
International Conference on the Applications of the Mössbauer Effect
Saint Petersburg, Russia, 3–8 September 2017

---
Spectrum
57Co/Rh source, α-iron (25μm) absorber
Theory
H, Is, (texture)
χ2 value
(close to expectation)?
Line width/shape?
area/thickness =>25μm ?
7 points concerning the
Spectrometer/Adaption of theory to experiment

---
Outline
Reduced χ2
Baseline of a transmission spectrum
The raw data problem
Choice of the drive frequency
Theoretical value at channel i
Dead time
Background
Conclusion: some good practices

---
Reduced χ2
size 12{ +- ~ sqrt { { {2} over {N rSub { size 8{ ital "ch"} } } } } ,} {}
χ2 =1.0
size 12{N rSub { size 8{ ital "ch"} } ="2048"~"=>"~χ rSup { size 8{2} } =1 +- 0 "." "031"} {}
Counts: Poisson distribution:

---
Reduced χ2
size 12{ +- ~ sqrt { { {2} over {N rSub { size 8{ ital "ch"} } } } } ,} {}
χ2 =1.0
size 12{N rSub { size 8{ ital "ch"} } ="2048"~"=>"~χ rSup { size 8{2} } =1 +- 0 "." "031"} {}
Counts: Poisson distribution:
size 12{ ital "Sp" left (v right ) prop Int cSub { size 8{ - infinity } } cSup { size 8{ infinity } } {L rSub { size 8{S} } } left (τ - v right )e rSup { size 8{ left ( - t rSub { ital "eff"} L rSub { size 8{A} } left (τ right ) right )} } ital "dτ"} {}
Beer–Lambert law
teff =10

---
Reduced χ2
size 12{ +- ~ sqrt { { {2} over {N rSub { size 8{ ital "ch"} } } } } ,} {}
χ2 =1.0
size 12{N rSub { size 8{ ital "ch"} } ="2048"~"=>"~χ rSup { size 8{2} } =1 +- 0 "." "031"} {}
Counts: Poisson distribution:
Fit: thin absorber approximation
size 12{ ital "Sp" left (v right ) prop 1 - t rSub { size 8{ ital "eff"} } L left ( ital "v,Γ=Γ" rSub { size 8{S} } ital "+Γ" rSub { size 8{A} } right )} {}
size 12{ ital "Sp" left (v right ) prop Int cSub { size 8{ - infinity } } cSup { size 8{ infinity } } {L rSub { size 8{S} } } left (τ - v right )e rSup { size 8{ left ( - t rSub { ital "eff"} L rSub { size 8{A} } left (τ right ) right )} } ital "dτ"} {}
Beer–Lambert law
teff =10

---
Reduced χ2
size 12{ +- ~ sqrt { { {2} over {N rSub { size 8{ ital "ch"} } } } } ,} {}
χ2 =1.0
size 12{N rSub { size 8{ ital "ch"} } ="2048"~"=>"~χ rSup { size 8{2} } =1 +- 0 "." "031"} {}
Counts: Poisson distribution:
size 12{ ital "Sp" left (v right ) prop 1 - t rSub { size 8{ ital "eff"} } L left ( ital "v,Γ=Γ" rSub { size 8{S} } ital "+Γ" rSub { size 8{A} } right )} {}
size 12{ ital "Sp" left (v right ) prop Int cSub { size 8{ - infinity } } cSup { size 8{ infinity } } {L rSub { size 8{S} } } left (τ - v right )e rSup { size 8{ left ( - t rSub { ital "eff"} L rSub { size 8{A} } left (τ right ) right )} } ital "dτ"} {}
Beer–Lambert law
teff =10
(AS70 algorithm of Odeh and Evans, 1974)
synthetic data by Monte Carlo simulations
Fit: thin absorber approximation
size 12{χ rSup { size 8{2} } =5 "." "47,"~ ital "Γ="2 "." "49"Γ rSub { size 8{N} } } {} size 12{χ rSup { size 8{2} } =1 "." "42,"~ ital "Γ="2 "." "15"Γ rSub { size 8{N} } ,~σ rSub { size 8{ ital "Gauss"} } =1 "." "17"Γ rSub { size 8{N} } } {}
Lorentz-curve:
Voigt profile:

---
Reduced χ2
Note:
Fit: thin absorber approximation
area is obtained to a good approximation
by the χ2 – fit procedure
Fit: finite thickness
Use of Beer–Lambert law and convolution integral: thickness instead of area χ2-value gets a meaning (validity of the theory)

---
Baseline of a transmission spectrum
Geometry
L0
detector
S0
+Smax
-Smax
aperture
Ra

---
Baseline of a transmission spectrum
Geometry
Undefined geometry
L0
detector
S0
+Smax
-Smax
aperture
Ra
detector
L0
aperture
Ra'

---
Baseline of a transmission spectrum
size 12{S rSub { size 8{v} } } {}
detector
L0
Ra

---
Baseline of a transmission spectrum
size 12{ %OMEGA left (v right ) ital "= "2π left (1 - { {L rSub { size 8{0} } ital "+S" left (v right )} over { sqrt { left (L rSub { size 8{0} } ital "+S" left (v right ) right ) rSup { size 8{ ital " 2"} } ital "+R" rSub { size 8{a} } rSup { size 8{ ital " 2"} } } } } right )} {} size 12{ %OMEGA rSub { size 8{0} } ital "= "2π~ left (1 - { {L rSub { size 8{0} } } over { sqrt {L rSub { size 8{0} } rSup { size 8{ ital " 2"} } ital "+R" rSub { size 8{a} } rSup { size 8{ ital " 2"} } } } } right ),} {}
Solid angle:
size 12{S left (v right )} {}
detector
L0
Ra

---
Baseline of a transmission spectrum
size 12{ { { %OMEGA left (v right )} over { %OMEGA rSub { size 8{0} } } } ital "= " { {1} over { left (1 ital "+δ" right ) rSup { size 8{2} } } } ~ left (1+ { {3} over {4} } ε rSup { size 8{2} } left (1 - { {1} over { left (1 ital "+δ" right ) rSup { size 8{2} } } } right )+ "." "." "." right )~ approx ~ { {1} over { left (1 ital "+δ" right ) rSup { size 8{2} } } } ,} {} size 12{δ left (v right ) ital "= " { {S left (v right )} over {L rSub { size 8{0} } } } } {} size 12{ ital "ε= " { {R rSub { size 8{a} } } over {L rSub { size 8{0} } } } ,} {}
Count rate (v):
size 12{S left (v right )} {}
detector
L0
Ra
size 12{ %OMEGA left (v right ) ital "= "2π left (1 - { {L rSub { size 8{0} } ital "+S" left (v right )} over { sqrt { left (L rSub { size 8{0} } ital "+S" left (v right ) right ) rSup { size 8{ ital " 2"} } ital "+R" rSub { size 8{a} } rSup { size 8{ ital " 2"} } } } } right )} {} size 12{ %OMEGA rSub { size 8{0} } ital "= "2π~ left (1 - { {L rSub { size 8{0} } } over { sqrt {L rSub { size 8{0} } rSup { size 8{ ital " 2"} } ital "+R" rSub { size 8{a} } rSup { size 8{ ital " 2"} } } } } right ),} {}
Solid angle:

---
Baseline of a transmission spectrum
size 12{ { {S left (v right )} over {L rSub { size 8{0} } } } ital "= " +- { {S rSub { size 8{ ital "max"} } } over {L rSub { size 8{0} } } } sqrt {1 - left ( { {v} over {v rSub { size 8{ ital "max"} } } } right ) rSup { size 8{2} } } ,} {}
size 12{S left (t right ) ital "= S" rSub { size 8{ ital "max"} } "sin"` left ( ital "ωt" right ),} {}

size 12{S rSub { size 8{ ital "max"} } `ω"= "v rSub { size 8{ ital "max"} } ,} {}
size 12{v left (t right ) ital "= S" rSub { size 8{ ital "max"} } ω`"cos"` left ( ital "ωt" right )} {}
size 12{ { { %OMEGA left (v right )} over { %OMEGA rSub { size 8{0} } } } ital "= " { {1} over { left (1 ital "+δ" right ) rSup { size 8{2} } } } ~ left (1+ { {3} over {4} } ε rSup { size 8{2} } left (1 - { {1} over { left (1 ital "+δ" right ) rSup { size 8{2} } } } right )+ "." "." "." right )~ approx ~ { {1} over { left (1 ital "+δ" right ) rSup { size 8{2} } } } ,} {} size 12{δ left (v right ) ital "= " { {S left (v right )} over {L rSub { size 8{0} } } } } {} size 12{ ital "ε= " { {R rSub { size 8{a} } } over {L rSub { size 8{0} } } } ,} {}
Count rate (v):
size 12{S left (v right )} {}
detector
L0
Ra
size 12{ %OMEGA left (v right ) ital "= "2π left (1 - { {L rSub { size 8{0} } ital "+S" left (v right )} over { sqrt { left (L rSub { size 8{0} } ital "+S" left (v right ) right ) rSup { size 8{ ital " 2"} } ital "+R" rSub { size 8{a} } rSup { size 8{ ital " 2"} } } } } right )} {} size 12{ %OMEGA rSub { size 8{0} } ital "= "2π~ left (1 - { {L rSub { size 8{0} } } over { sqrt {L rSub { size 8{0} } rSup { size 8{ ital " 2"} } ital "+R" rSub { size 8{a} } rSup { size 8{ ital " 2"} } } } } right ),} {}
Solid angle:

---
Baseline of a transmission spectrum
size 12{ { { %OMEGA left (v right )} over { %OMEGA rSub { size 8{0} } } } ital "= " { {1} over { left (1 ital "+δ" right ) rSup { size 8{2} } } } ~ left (1+ { {3} over {4} } ε rSup { size 8{2} } left (1 - { {1} over { left (1 ital "+δ" right ) rSup { size 8{2} } } } right )+ "." "." "." right )~ approx ~ { {1} over { left (1 ital "+δ" right ) rSup { size 8{2} } } } ,} {} size 12{δ left (v right ) ital "= " { {S left (v right )} over {L rSub { size 8{0} } } } } {} size 12{ ital "ε= " { {R rSub { size 8{a} } } over {L rSub { size 8{0} } } } ,} {}
Count rate (v):
size 12{S left (t right ) ital "= S" rSub { size 8{ ital "max"} } "sin"` left ( ital "ωt" right ),} {} size 12{v left (t right ) ital "= S" rSub { size 8{ ital "max"} } ω`"cos"` left ( ital "ωt" right )} {} size 12{ { {S left (v right )} over {L rSub { size 8{0} } } } ital "= " +- { {S rSub { size 8{ ital "max"} } } over {L rSub { size 8{0} } } } sqrt {1 - left ( { {v} over {v rSub { size 8{ ital "max"} } } } right ) rSup { size 8{2} } } ,} {} size 12{S rSub { size 8{ ital "max"} } `ω"= "v rSub { size 8{ ital "max"} } ,} {} size 12{S rSub { size 8{ ital "max"} } ital "= " { {v rSub { size 8{ ital "max"} } } over {2 ital "πf"} } } {}

Sinusoidal:
size 12{S rSub { size 8{ ital "max"} } ital "= " { {v rSub { size 8{ ital "max"} } } over {8f} } } {}
Triangular:
size 12{S left (v right )} {}
detector
L0
Ra
size 12{ %OMEGA left (v right ) ital "= "2π left (1 - { {L rSub { size 8{0} } ital "+S" left (v right )} over { sqrt { left (L rSub { size 8{0} } ital "+S" left (v right ) right ) rSup { size 8{ ital " 2"} } ital "+R" rSub { size 8{a} } rSup { size 8{ ital " 2"} } } } } right )} {} size 12{ %OMEGA rSub { size 8{0} } ital "= "2π~ left (1 - { {L rSub { size 8{0} } } over { sqrt {L rSub { size 8{0} } rSup { size 8{ ital " 2"} } ital "+R" rSub { size 8{a} } rSup { size 8{ ital " 2"} } } } } right ),} {}
Solid angle:

---
Baseline of a transmission spectrum
Note:
The baseline is determined by three parameters:
size 12{ ital "geo=" { {S rSub { size 8{"max"} } } over {L rSub { size 8{0} } } } ,} {}
size 12{ ital "counts"` left ( ital "v=" infinity right )`,} {}
size 12{ ital "channel"` rSub { size 8{v`=`0} } } {}

---
Baseline of a transmission spectrum
Note:
The baseline is determined by three parameters:
size 12{ ital "geo=" { {S rSub { size 8{"max"} } } over {L rSub { size 8{0} } } } ,} {}
size 12{ ital "channel"` left ( ital "v="0 right )`} {}
size 12{ ital "counts"` left ( ital "v=" infinity right )`,} {}
size 12{ ital "counts"` left ( ital "v=" infinity right )`,} {}
( fit! )
geo = 5.77·10-4
size 12{ ital "geo= " { {v rSub { size 8{"max"} } } over {2 ital "πL" rSub { size 8{0} } f} } = { {7 "." "24"`"mm/s"} over {2π cdot "120"`"mm"` cdot "17"`"Hz"} } } {}
= 5.65·10-4
size 12{χ rSup { size 8{2} } "= 1" "." "007"} {}
(120 mm → 117.5 mm)

---
The raw data problem
sinusoidal
triangular
Calibration spectrum for FAS
W. C. Tennant, Christchurch (2009).

---
The raw data problem
Fit of 6 independent lines
(position, Voigt profile)
sinusoidal
triangular

---
The raw data problem
Fit of 6 independent lines
(position, Voigt profile)
sinusoidal
triangular

---
The raw data problem
Fit of 6 independent lines
(position, Voigt profile)
sinusoidal
triangular

---
The raw data problem
size 12{Z left ( ital "FP" right )= Sum cSub { size 8{ ital "i>FP"} } cSup { size 8{N} } { left (C rSub { size 8{i} } - C rSub { size 8{ ital "FP" - left (i - ital "FP" right )} } right ) rSup { size 8{2} } } } {} size 12{χ rSup { size 8{2} } =1 "." "61"} {}
Folding point FP: minimum of Z(FP)
FP=512.5
6 Voigt profiles

---
The raw data problem
size 12{Z left ( ital "FP" right )= Sum cSub { size 8{ ital "i>FP"} } cSup { size 8{N} } { left (C rSub { size 8{i} } - C rSub { size 8{ ital "FP" - left (i - ital "FP" right )} } right ) rSup { size 8{2} } } } {} size 12{χ rSup { size 8{2} } =1 "." "61"} {}
Folding point FP: minimum of Z(FP)
FP=512.5
6 Voigt profiles

---
The raw data problem
Correction of the nonlinearity of the velocity scale
v(i)
size 12{v left (i right )= - v rSub { size 8{ ital "max"} } { {4} over {π rSup { size 8{2} } } } left ("cos"ω cdot ital "i+" { {"cos3"ω cdot i} over {3 rSup { size 8{2} } } } + { {"cos5"ω cdot i} over {5 rSup { size 8{2} } } } + "." "." "." right )} {}
size 12{ω cdot i} {}
size 12{ ital "dv" left (i right )= Sum {a rSub { size 8{k} } } "cos" left ( left (2 ital "k+"1 right )ω cdot i right ) ital "+b" rSub { size 8{k} } "sin" left ( left (2 ital "k+"1 right )ω cdot i right )} {}
size 12{2π} {}
size 12{ω cdot "256" ital "=π"/2} {}
size 12{π} {}

---
The raw data problem
Correction of the nonlinearity of the velocity scale
v(i)
size 12{ω cdot i} {}
size 12{2π} {}
size 12{ω cdot "256" ital "=π"/2} {}
size 12{π} {}
size 12{χ rSup { size 8{2} } =1 "." "051"} {}
Source (5mCi):
ГLorentz = 1.03 ГN , бGauss=0.45 ГN
dv = ± 7.5 channels
size 12{Γ left (α"-iron" right ) ital "=Γ" rSub { size 8{N} } } {}
size 12{v left (i right )= - v rSub { size 8{ ital "max"} } { {4} over {π rSup { size 8{2} } } } left ("cos" ital "ωi+" { {"cos3"ω cdot i} over {3 rSup { size 8{2} } } } + { {"cos5"ω cdot i} over {5 rSup { size 8{2} } } } + "." "." "." right )} {} size 12{ ital "dv" left (i right )= Sum {a rSub { size 8{k} } } "cos" left ( left (2 ital "k+"1 right )ω cdot i right ) ital "+b" rSub { size 8{k} } "sin" left ( left (2 ital "k+"1 right )ω cdot i right )} {}

---
The raw data problem
Not only positions of the lines but also their shapes
determine the velocity scale
Note:
No reliable velocity scale after folding!!
Better solution by a fit of the full spectrum and
velocity correction for each channel

---
Choice of the drive frequency
17.0 Hz
4.25 Hz
chv=0=513.08
triangular
size 12{ ital "geo=" { {v rSub { size 8{ ital "max"} } } over {8L rSub { size 8{0} } f} } } {}
J Pechousek et al., Palacky University of Olomouc, Czech Republic
www.researchgate.net/publication/252974211
Vmax=7.24 mm/s
L0=117.5 mm
size 12{χ rSup { size 8{2} } =1 "." "028"} {} size 12{χ rSup { size 8{2} } =1 "." "035"} {}

---
Choice of the drive frequency
17.0 Hz
4.25 Hz
chv=0=513.08
triangular
size 12{ ital "geo=" { {v rSub { size 8{ ital "max"} } } over {8L rSub { size 8{0} } f} } } {}
size 12{ ital "geo=" { {v rSub { size 8{ ital "max"} } } over {2 ital "πL" rSub { size 8{0} } f} } } {}
sinusiodal
J Pechousek et al., Palacky University of Olomouc, Czech Republic
www.researchgate.net/publication/252974211
Vmax=7.24 mm/s
L0=117.5 mm
size 12{χ rSup { size 8{2} } =1 "." "028"} {} size 12{χ rSup { size 8{2} } =1 "." "035"} {}
size 12{χ rSup { size 8{2} } =1 "." "027"} {} size 12{χ rSup { size 8{2} } =1 "." "084"} {}

---
Choice of the drive frequency
Note:
Nonlinearity correction dv(i) strongly depends on the drive frequency
The evaluated physical parameters of the spectrum
shall not depend on the driving mode, the frequency nor the solid angle

---
Theoretical value at channel i
size 12{ ital "Th" rSub { size 8{i} } = Int cSub { size 8{t rSub {i} } } cSup {t rSub { size 8{ ital "i+"1} } } { ital "Th"} left (v left (t right ) right ) ital "dt"} {}

---
Theoretical value at channel i
size 12{ ital "Th" rSub { size 8{i} } = ital "Δt" cdot ital "Th" left (v rSub { size 8{i} } right )} {}
size 12{ ital "Th" rSub { size 8{i} } = Int cSub { size 8{t rSub {i} } } cSup {t rSub { size 8{ ital "i+"1} } } { ital "Th"} left (v left (t right ) right ) ital "dt"} {}

---
Theoretical value at channel i
size 12{ ital "Th" rSub { size 8{i} } = ital "Δt" cdot ital "Th" left (v rSub { size 8{i} } right )} {}
size 12{ ital "Th" rSub { size 8{i} } = Int cSub { size 8{t rSub {i} } } cSup {t rSub { size 8{ ital "i+"1} } } { ital "Th"} left (v left (t right ) right ) ital "dt"} {}
size 12{ ital "Th" rSub { size 8{i} } = Int cSub { size 8{t rSub {i} } } cSup {t rSub { size 8{ ital "i+"1} } } { ital "Th"} left (v left (t right ) right ) { { ital "dt"} over { ital "dv"} } ital "dv"} {} size 12{ {}= { { ital "Δt"} over {v left (t rSub { size 8{ ital "i+"1} } right ) - v left (t rSub { size 8{i} } right )} } Int cSub { size 8{v left (t rSub {i} right )} } cSup {v left (t rSub { size 8{ ital "i+"1} } right )} { ital "Th"} left (v right ) ital "dv"} {}

---
Theoretical value at channel i
best parabola by 4 points

---
Theoretical value at channel i
best parabola by 4 points

---
Theoretical value at channel i
Note:
Advantages of theory interpolation:
Fit of dvi dependent on 6-18 parameters would not be feasible
with the convolution for each iteration step: 1024-4096 numerical integrals

---
Theoretical value at channel i
Note:
Calculation of cos-smearing for large Ω can be easily added (weak sources)
aperture
Weighted (by Ωk) superposition of spectra with vi cos(ϑk )
Advantages of theory interpolation:
Fit of dvi dependent on 6-18 parameters would not be feasible
with the convolution for each iteration step: 1024-4096 numerical integrals
Ωk
ϑk

---
Dead time
paralyzable, nonparalyzable
Proportional counter
size 12{ ital "τ=""100" ital "μs"} {}
size 12{ ital "τ="1 ital "μs"} {}
Scintillator
D.J. Morrissey,2oo9, Michigan State University

---
Dead time
paralyzable, nonparalyzable
Proportional counter
size 12{ ital "τ=""100" ital "μs"} {}
size 12{ ital "τ="1 ital "μs"} {}
Scintilator
size 12{N rSub { size 8{ ital "obs"} } = { {N} over {1 ital "+Nτ" rSub { size 8{ ital "np"} } } } } {}
size 12{N rSub { size 8{ ital "obs"} } = { { ital "Ne" rSup { size 8{ - ital "Nτ" rSub {p} } } } over {1 ital "+Nτ" rSub { size 8{ ital "np"} } } } } {}
size 12{N rSub { size 8{ ital "obs"} } ital "=Ne" rSup { size 8{ - ital "Nτ" rSub {p} } } } {}
D.J. Morrissey,2oo9, Michigan State University
N
Nobs

---
Dead time
Code for Monte Carlo simulations (nonparalyzable, paralyzable)
while(isimul[0] < *counts)
}
eventtime=-log(1.-frand())/(countrate*theo[ichannel]);
lastevent=lastevent+eventtime;
totaltime=totaltime+eventtime;
ichannel= totaltime/dwelltime;
i=ichannel/nu_channel; ichannel=ichannel-i*nu_channel;
if(lastevent > deadtime)
{isimul[ichannel]++; lastevent=0.0; icount++;}
else
{iloss++;}
{iloss++;lastevent=0.;}
}

---
Dead time
Nonparalyzable
paralyzable

---
Dead time
Nonparalyzable
paralyzable

---
Dead time
Nonparalyzable
paralyzable

---
Dead time
paralyzable
N
Nobs
size 12{ ital "Nτ" <= 0 "." 5} {}
up to
deadtime effects are kept within reasonable limits

---
Dead time
size 12{ ital "geo= " { {v rSub { size 8{ ital "max"} } } over {2 ital "πL" rSub { size 8{0} } f} } = { {"15" "." 6 ital "mm"/s} over {2π cdot bold "67" ital "mm" cdot "16" ital "Hz"} } " =2" "." "32" cdot "10" rSup { size 8{ - 3} } } {}
size 12{χ rSup { size 8{2} } =0 "." bold "994"} {}
5.5mCi
size 12{ ital "N=counts" left (v rSub { size 8{ infinity } } right ) { { %OMEGA left ( ital "geo,v" rSub { size 8{i} } right )} over { %OMEGA rSub { size 8{0} } } } } {}
(fit!)
size 12{" => " ital "geo" bold "= 1" "." bold "62" cdot bold "10" rSup { size 8{ - 3} } } {}

---
Dead time
size 12{ ital "geo= " { {v rSub { size 8{ ital "max"} } } over {2 ital "πL" rSub { size 8{0} } f} } = { {"15" "." 6 ital "mm"/s} over {2π cdot bold "67" ital "mm" cdot "16" ital "Hz"} } " =2" "." "32" cdot "10" rSup { size 8{ - 3} } } {}
size 12{χ rSup { size 8{2} } =0 "." bold "994"} {}
5.5mCi
size 12{ ital "N=counts" left (v rSub { size 8{ infinity } } right ) { { %OMEGA left ( ital "geo,v" rSub { size 8{i} } right )} over { %OMEGA rSub { size 8{0} } } } } {}
(fit!)
size 12{" => " ital "geo" bold "= 1" "." bold "62" cdot bold "10" rSup { size 8{ - 3} } } {}
size 12{,~N rSub { size 8{c} } = { {N} over {1 ital "+Nτ" rSub { size 8{ ital "np"} } } } } {}
size 12{ ital "N=counts" left (v rSub { size 8{ infinity } } right ) { { %OMEGA left (2 "." "32" cdot "10" rSup { size 8{ - 3} } ital ",v" rSub { size 8{i} } right )} over { %OMEGA rSub { size 8{0} } } } } {}
(fit!)
size 12{" => " ital "Nτ" rSub { size 8{ ital "np"} } =0 "." bold "42"} {}

---
Dead time
size 12{ ital "geo= " { {v rSub { size 8{ ital "max"} } } over {2 ital "πL" rSub { size 8{0} } f} } = { {"15" "." 6 ital "mm"/s} over {2π cdot bold "67" ital "mm" cdot "16" ital "Hz"} } " =2" "." "32" cdot "10" rSup { size 8{ - 3} } } {}
size 12{χ rSup { size 8{2} } =0 "." bold "994"} {}
5.5mCi
size 12{ ital "N=counts" left (v rSub { size 8{ infinity } } right ) { { %OMEGA left ( ital "geo,v" rSub { size 8{i} } right )} over { %OMEGA rSub { size 8{0} } } } } {}
(fit!)
size 12{" => " ital "geo" bold "= 1" "." bold "62" cdot bold "10" rSup { size 8{ - 3} } } {}
size 12{,~N rSub { size 8{c} } = { {N} over {1 ital "+Nτ" rSub { size 8{ ital "np"} } } } } {}
size 12{ ital "N=counts" left (v rSub { size 8{ infinity } } right ) { { %OMEGA left (2 "." "32" cdot "10" rSup { size 8{ - 3} } ital ",v" rSub { size 8{i} } right )} over { %OMEGA rSub { size 8{0} } } } } {}
(fit!)
size 12{" => " ital "Nτ" rSub { size 8{ ital "np"} } =0 "." bold "42"} {}
size 12{ { {2 "." "32"} over { left (1 ital "+Nτ" rSub { size 8{ ital "np"} } right )} } =1 "." "63"} {}

---
Dead time
122/136keV
14.4keV
NaI
(ungated)
size 12{N rSub { size 8{c} } = { {N} over {1 ital "+Nτ" rSub { size 8{ ital "np"} } } } } {}
nonparalyzable

---
Dead time
122/136keV
14.4keV
NaI
(ungated)
size 12{ ital "Nτ" rSub { size 8{ ital "np"} } "=0" "." "22"~ rightarrow ~τ rSub { size 8{ ital "np"} } =4 "." 4 ital "μs"} {}
size 12{N rSub { size 8{c} } = { {N} over {1 ital "+Nτ" rSub { size 8{ ital "np"} } } } } {}
nonparalyzable

---
Dead time
122/136keV
14.4keV
(gated)
NaI
size 12{ ital "Nτ" rSub { size 8{ ital "np"} } "=0" "." "22"~ rightarrow ~τ rSub { size 8{ ital "np"} } <4 "." 4 ital "μs"} {}
size 12{N rSub { size 8{c} } = { {N} over {1 ital "+Nτ" rSub { size 8{ ital "np"} } } } } {}
nonparalyzable

---
Dead time
122/136keV
14.4keV
(gated)
NaI
size 12{ ital "Nτ" rSub { size 8{ ital "np"} } "=0" "." "22"} {}
size 12{N rSub { size 8{c} } = { {N} over {1 ital "+Nτ" rSub { size 8{ ital "np"} } } } } {}
nonparalyzable
d=114mm
15mmØ
2mCi

---
Dead time
122/136keV
14.4keV
(gated)
NaI
size 12{ ital "Nτ" rSub { size 8{ ital "np"} } "=0" "." "22"} {}
size 12{N rSub { size 8{c} } = { {N} over {1 ital "+Nτ" rSub { size 8{ ital "np"} } } } } {}
nonparalyzable
d=114mm
15mmØ
2mCi
50mCi → 5d

---
Dead time
Dead time
Note:
Avoid (severe) dead time effects !
Check by the distance law 1/r2 for dead time
Check by the geo parameter for dead time:
up to
deadtime effects are kept within reasonable limits
size 12{ ital "Nτ" <= 0 "." 5} {}
size 12{ { { ital "geo"} over { ital "geo" rSub { size 8{ ital "eff"} } } } = left (1 ital "+Nτ" right )} {}

---
Background
6.4 keV
Amptek (CdTe)
1.8keV
Kr
14.4
6.4
~22
A straightforward experimental method to evaluate the Lamb-Mossbauer factor of a 57Co/Rh source
G. Spina , M. Lantieri, Nul. Instrum. Methods B 318 (2014) 253-257
detector

---
Background
6.4 keV
Amptek (CdTe)
detector
100μm Ni-foil
14.4
1.8keV
6.4
~22
Kr

---
Background
6.4 keV
Ni-foil
Position behind source
Amptek (CdTe)
Amptek (CdTe)
detector
14.4
1.8keV
6.4
~22
Kr

---
Background
6.4 keV
Position behind source
Amptek (CdTe)
Amptek (CdTe)
size 12{ ital "bg" rSub { size 8{ ital "fr"} } = { {A rSub { size 8{b} } \( bold "Ni" \) /0 "." bold "9695"} over {A rSub { size 8{t} } } } } {}
14.4keV - window
multiscaling mode
size 12{A rSub { size 8{b} } \( bold "Ni" \) ital ",A" rSub { size 8{t} } } {}
detector
Ni-foil
14.4
1.8keV
6.4
~21
Kr
size 12{"110"μ"Ni: 0" "." "9695"} {}
Attenuation of 122/136 keV

---
Background
size 12{b rSub { size 8{ ital "fr"} } = { {A rSub { size 8{b} } \( bold "Ni" \) /0 "." bold "9695"} over {A rSub { size 8{t} } /0 "." bold "9877"} } =1 "." bold "019" { {A rSub { size 8{b} } \( bold "Ni" \) } over {A rSub { size 8{t} } } } } {}
1.23%
A straightforward experimental method to evaluate the Lamb-Mossbauer factor of a 57Co/Rh source
G. Spina , M. Lantieri, Nul. Instrum. Methods B 318 (2014) 253-257
1.23%

---
Background
size 12{b rSub { size 8{ ital "fr"} } = { {A rSub { size 8{b} } \( bold "Ni" \) /0 "." bold "9695"} over {A rSub { size 8{t} } /0 "." bold "9877"} } =1 "." bold "019" { {A rSub { size 8{b} } \( bold "Ni" \) } over {A rSub { size 8{t} } } } } {}
1.23%
Green: 33.05 Tesla
size 12{α - ital "iron"~"25" "." 5 ital "μm,"~ ital "f="0 "." "80,"~ ital "Γ=Γ" rSub { size 8{N} } } {}
Red: 30.67 Tesla (Mn)
size 12{χ rSup { size 8{2} } =0 "." "984"} {}
Mossbauer measurements in iron based alloys with transition metals I. Vincze and A. Campbell,J. Phys. F: Metal Phys. Vol.3 (1973) 647-663
1.23%

---
Background
size 12{b rSub { size 8{ ital "fr"} } = { {A rSub { size 8{b} } \( bold "Ni" \) /0 "." bold "9695"} over {A rSub { size 8{t} } /0 "." bold "9877"} } =1 "." bold "019" { {A rSub { size 8{b} } \( bold "Ni" \) } over {A rSub { size 8{t} } } } } {}
1.23%
Green: 33.05 Tesla
size 12{α - ital "iron"~"25" "." 5 ital "μm,"~ ital "f="0 "." "80,"~ ital "Γ=Γ" rSub { size 8{N} } } {}
Red: 30.67 Tesla (Mn)
size 12{χ rSup { size 8{2} } =0 "." "984"} {}
Mossbauer measurements in iron based alloys with transition metals I. Vincze and A. Campbell,J. Phys. F: Metal Phys. Vol.3 (1973) 647-663
1.23%
size 12{Γ rSub { size 8{ ital "source"} } =1 "." "38" cdot Γ rSub { size 8{N} } } {} size 12{σ rSub { size 8{ ital "source"} } =0 "." "59" cdot Γ rSub { size 8{N} } } {}
fsource(1-bgfr) =>
fsource = 0.715
Fit results:

---
Conclusion: some good practices
Taking down to the logbook:
Source (link to data sheet), Absorber (mg/cm2), Geometry (L0,aperture),
drive(mode, frequency), counting system (detector, electronic settings)
Background fraction (rate measurements: gated/ungated, Ni-foil), etc
Use of the convolution integral
Use of a reliable nonlinearity correction
-Reliable values of teff instead relative areas
(independent of the complexity of the applied theory)
-Continuous control of the components of the equipment
Advantage: